SAP field calibration instructions

It is possible to recalibrate the SAP in the field. It does not require a compass or clino. It is actually easier to do underground than on the surface; it is important that there are no ferrous objects nearby. This includes your watch.

1. Turn the device on, putting it in ”Daylight Mode” by holding down the button on the second click
2. Press the button twice in quick succession.
3. The display should now read C000 - you are now in field calibration mode.
4. Choose two fixed points, lets call them A and B.
5. Take four readings from A to B, with the display facing up, left, down and right respectively.
6. The display should now read C004
7. Repeat from B to A, the display should now read C008
8. Choose a further two fixed points, C and D so that C->D is about 90 degrees from A->B
9. Take four readings from C to D and four from D to C.
10. The display should now read C008
11. We now need to collect a further 48 legs of random data. Imagine A,B,C,D being in the middle of four sides of a cube. Take a set of four readings from each of the corners of the cube, and one set straight up and one set straight down.
12. The display should start slowly counting up from CLX1 to CL99 - it will take about 5-10 minutes.
13. Once complete, turn the device off by pressing the button briefly.
14. The calibration process is now complete. It can be cancelled at any point up to step 12 by turning the device off